## Moorings
Mooring OOIBARCODE | Ref Des | Serial Number | Deployment Number | Anchor Launch Date | Anchor Launch Time | Recover Date | Latitude | Longitude | Water Depth | Cruise Number | Notes
ATAPL-71403-00001 | RS03AXPD-DP03A | SN0001 | 1 | 2014-08-09 | 21:00 |  | 45° 49.7832'N | 129° 45.5424'W | 2612 | TN-313 | Profiler no longer operating due to docking station overcurrent | 45.82972 | -129.75904
ATAPL-67977-00002 | RS03AXPD-DP03A-01-CTDPFL304 | 5277187-0137 | 1 | 2014-08-09 | 21:00 |  | 45° 49.7832'N | 129° 45.5424'W | 2612 | TN-313 | Profiler not operating due to docking station overcurrent | 45.82972 | -129.75904
ATAPL-58346-00002 | RS03AXPD-DP03A-02-VEL3DA303 | 1130 | 1 | 2014-08-09 | 21:00 |  | 45° 49.7832'N | 129° 45.5424'W | 2612 | TN-313 | Profiler not operating due to docking station overcurrent | 45.82972 | -129.75904
ATAPL-70111-00002 | RS03AXPD-DP03A-03-FLCDRA302 | 3399 | 1 | 2014-08-09 | 21:00 |  | 45° 49.7832'N | 129° 45.5424'W | 2612 | TN-313 | Profiler no longer operating due to docking station overcurrent | 45.82972 | -129.75904
ATAPL-70110-00002 | RS03AXPD-DP03A-04-FLNTUA302 | 3397 | 1 | 2014-08-09 | 21:00 |  | 45° 49.7832'N | 129° 45.5424'W | 2612 | TN-313 | Profiler no longer operating due to docking station overcurrent | 45.82972 | -129.75904
ATAPL-58320-00001 | RS03AXPD-DP03A-06-DOSTAD304 | 458 | 1 | 2014-08-09 | 21:00 |  | 45° 49.7832'N | 129° 45.5424'W | 2612 | TN-313 | Profiler no longer operating due to docking station overcurrent | 45.82972 | -129.75904
ATAPL-71553-00005 | RS03AXPD-PD03A | SN0005 | 1 | 2014-08-09 | 21:00 |  | 45° 49.7832'N | 129° 45.5424'W | 2612 | TN-313 | Overcurrent, likely due to known dummy plug failure Docking Station no longer powered. | 45.82972 | -129.75904

## Asset_Cal_Info
Ref Des | Mooring OOIBARCODE | Mooring Serial Number | Deployment Number | Sensor OOIBARCODE | Sensor Serial Number | Calibration Cofficient Name | Calibration Cofficient Value | Notes
RS03AXPD-DP03A-01-CTDPFL304 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-67977-00002 | 5277187-0137 | CC_lat | 45.8297 | Depth
RS03AXPD-DP03A-01-CTDPFL304 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-67977-00002 | 5277187-0137 | CC_lon | -129.759
Compute the L2 PRACSAL data product from CTD
RS03AXPD-DP03A-02-VEL3DA303 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-58346-00002 | 1130 | CC_lat | 45.8297 | Depth
RS03AXPD-DP03A-02-VEL3DA303 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-58346-00002 | 1130 | CC_lon | -129.759
RS03AXPD-DP03A-02-VEL3DA303 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-58346-00002 | 1130 | CC_hdg_cal | [335.0 ,20.0 ,65.0 ,110.0 ,155.0 ,200.0 ,245.0 ,290.0 ]
RS03AXPD-DP03A-02-VEL3DA303 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-58346-00002 | 1130 | CC_hx_cal | [0.4,0.44,0.25,-0.04,-0.28,-0.31,-0.1,0.19]
RS03AXPD-DP03A-02-VEL3DA303 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-58346-00002 | 1130 | CC_hy_cal | [-0.1,0.18,0.42,0.47,0.29,-0.01,-0.26,-0.29]
RS03AXPD-DP03A-03-FLCDRA302 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-70111-00002 | 3399 | CC_scale_factor_cdom | 0.0301 | Depth
RS03AXPD-DP03A-03-FLCDRA302 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-70111-00002 | 3399 | CC_dark_counts_cdom | 46
RS03AXPD-DP03A-04-FLNTUA302 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-70110-00002 | 3397 | CC_scale_factor_volume_scatter | 0 | Not Available. Calibrated for Nephelometric Turbidity.
RS03AXPD-DP03A-04-FLNTUA302 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-70110-00002 | 3397 | CC_dark_counts_volume_scatter | 0 | Not Available. Calibrated for Nephelometric Turbidity.
RS03AXPD-DP03A-04-FLNTUA302 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-70110-00002 | 3397 | CC_scale_factor_chlorophyll_a | 0.0121
RS03AXPD-DP03A-04-FLNTUA302 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-70110-00002 | 3397 | CC_dark_counts_chlorophyll_a | 49
RS03AXPD-DP03A-04-FLNTUA302 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-70110-00002 | 3397 | CC_scattering_angle | 117 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
RS03AXPD-DP03A-04-FLNTUA302 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-70110-00002 | 3397 | CC_measurement_wavelength | 700 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
RS03AXPD-DP03A-04-FLNTUA302 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-70110-00002 | 3397 | CC_angular_resolution | 1.08 | Default value per <flo_bback_total(beta, degC=20.0, psu=32.0, theta=117.0, wlngth=700.0, xfactor=1.08)>
RS03AXPD-DP03A-04-FLNTUA302 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-70110-00002 | 3397 | CC_depolarization_ratio | 0.039 | Default value per <flo_scat_seawater(degC, psu, theta=117.0, wlngth=700.0, delta=0.039)>
RS03AXPD-DP03A-06-DOSTAD304 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-58320-00001 | 458 | CC_lat | 45.8297 | Depth
RS03AXPD-DP03A-06-DOSTAD304 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-58320-00001 | 458 | CC_lon | -129.759 | Requires TEMPWAT, PRESWAT, and PRACSAL from RS03AXPD-DP03A-01-CTDPFL304
RS03AXPD-DP03A-06-DOSTAD304 | ATAPL-71403-00001 | SN0001 | 1 | ATAPL-58320-00001 | 458 | CC_csv | [2.72962e-03, 1.14611e-04, 2.21263e-06 ,2.33768e02, -2.94251e-01, -5.33864e01, 4.55129e00]
RS03AXPD-DP03A-05-OPTAAC302 was not deployed.

## IntegrationEvents
OOIBARCODE | Int_Asset | DESCRIPTION | Type | serial_number | Date | comments
ATAPL-71403-00001 | RS03AXPD-DP03A |  | Mooring | SN0001 | 8/9/2014
ATAPL-67977-00002 |  | RS03AXPD-DP03A-01-CTDPFL304 | Sensor | 5277187-0137
ATAPL-58346-00002 |  | RS03AXPD-DP03A-02-VEL3DA303 | Sensor | 1130
ATAPL-70111-00002 |  | RS03AXPD-DP03A-03-FLCDRA302 | Sensor | 3399
ATAPL-70110-00002 |  | RS03AXPD-DP03A-04-FLNTUA302 | Sensor | 3397
ATAPL-58320-00001 |  | RS03AXPD-DP03A-06-DOSTAD304 | Sensor | 458
ATAPL-71553-00005 | RS03AXPD-PD03A |  | Mooring | SN0005 | 8/9/2014

## Verification
Ref Des | uft21 load | Science Map (name) | Deployment | Calibration | Plot | production load | Science Map (name) | Plot
RS03AXPD-DP03A |  | no | 1/1
RS03AXPD-DP03A-01-CTDPFL304 |  | no | 1/1 | 1/1
RS03AXPD-DP03A-02-VEL3DA303 |  | no | 1/1 | 1/1
RS03AXPD-DP03A-03-FLCDRA302 |  | no | 1/1 | 1/1
RS03AXPD-DP03A-04-FLNTUA302 |  | no | 1/1 | 1/1
RS03AXPD-DP03A-06-DOSTAD304 |  | no | 1/1 | 1/1
RS03AXPD-PD03A |  | no | 1/1
| 0/0 | '0/7 | '7/7 | '5/5